Sr. No. in Scope NABL / NON NABL

 **I-B: Flow chart for analysis of ---------- AI Content in formulation sample**

|  |  |
| --- | --- |
| **Date of Analysis**  |  |

|  |  |  |  |
| --- | --- | --- | --- |
| **Sl. No.** | **Step**  | **Execution** | **Executed By**  |
| 1. | Sample No. |  |  |
| 2. | Name of Sample |
| 3. | **Procedure** |
| **3.1** | **Preparation of Mobile Phase** |  |  |
| 3.1.1 | Mix ----------- and -------- in the proportion of ----:---- (v/v)  |   |  |
| 3.1.2 | Pass through membrane filter under vacuum |  |  |
| 3.1.3 | Homogenize the mixture using a magnetic stirrer |  |  |
| 3.1.4 | Allow to attain room temperature & degas |  |  |
| **3.2** | **Preparation of Internal Standard Solution** |  |  |
| 3.2.1 | Weigh ---- g of -----Internal Standard in ---- ml volumetric flask. |  g |  |
| 3.2.2 | *Note the serial No. of the balance log book* |  |  |
| 3.2.3 | Add to it -- ml of mobile phase (3.1.4) or suitable solvent |  |  |
| 3.2.4 | Keep it for 30 minute with intermittent shaking |  |  |
| 3.2.5 | Make up to the mark with mobile phase (3.1.4) |   |  |
| **3.3**  | **Preparation of Standard Solution** |  |  |
| 3.3.1 | Note the purity of the standard | % |  |
| 3.3.2 | Weigh ---- g A.I. of standard in ---- ml volumetric flask  | g |  |
| 3.3.3 | *Note the serial No. of the balance log book* |  |  |
| 3.3.4 | Add to it --- ml of mobile phase (3.1.4) or suitable solvent |  |  |
| 3.3.5 | Keep it for 30 minute with intermittent shaking |  |  |
| 3.3.6 | Make up to the mark with mobile phase (3.1.4) |  |  |
| 3.3.7 | Pipette out--- ml of solution (3.3.6) to a --- ml volumetric flask if required  |  ml |  |
| 3.3.8 | Make up to the mark with mobile phase (3.1.4) or suitable solvent |  |  |
| 3.3.9 | Pipette out --- ml of solution (3.3.8) to a --- ml volumetric flask | ml |  |
| 3.3.10 | Add --ml of internal standard solution (3.2.5) | ml |  |
| 3.3.11 | make up to the mark with mobile phase (3.1.4) |  |  |
| **3.4** | **Preparation of Sample** |  |  |
| 3.4.1 | Note the percent active ingredient content declared on sample | % |  |
| 3.4.2 | Weigh -- g A.I. of sample in --- ml volumetric flask | g |  |
| 3.4.3 | *Note the serial No. of the balance log book* |  |  |
| 3.4.4 | Add to it --- ml of mobile phase (3.1.4) |  |  |
| 3.4.5 | Keep it for 30 minute with intermittent shaking |  |  |
| 3.4.6 | Male up to the mark with mobile phase (3.1.4) |  |  |
| 3.4.7 | Pipette out --- ml of solution (3.4.6) to a --- ml volumetric flask |  ml |  |
| 3.4.8 | Make up to the mark with mobile phase (3.1.4) |  |  |
| 3.4.9 | Pipette out --ml of solution (3.4.8) to a -- ml volumetric flask | ml |  |
| 3.4.10 | Add --ml of internal standard solution (3.2.5) | ml |  |
| 3.4.11 | Make up to the mark with mobile phase (3.1.4) |  |  |
| **4.** | **HPLC Parameters** |  |  |
| **4.1** | **Column** |  |  |
| 4.1.1 | Stainless Steel Packed with ----- Stationary phase |  |  |
| 4.1.2 | Length: ------ mm |  |  |
| 4.1.3 | I.D.: -------mm |  |  |
| **4.2** | **Mobile Phase** |  |  |
| 4.2.1 | --------- : ------- (-- : --) |  |  |
| 4.2.2 | Flow Rate: ----- ml/min |  |  |
| **4.3** | **Detector:** UV |  |  |
| **4.4**  | **Wave Length**: ---- nm |  |  |
| **4.5** | **Injection Volume: ----** µl |  |  |
| **5.** | **Result** |  |  |
| Sample chromatogram no.  |  |
| Standard chromatogram no.  |  |

**6. Calculation:**

 A2 x A3 x M1

 Imidacloprid content, % by mass = ------------------- x P

 A1 x A4 x M2

**Where,**

M1 = Mass in ‘g’ of ---- Pesticide standard

 M2 =Mass in ‘g’ of sample taken for test

 A1 = Peak area of Pesticide AI \_\_\_\_\_\_ in the standard solution

 A2 = Peak area of Pesticide AI \_\_\_\_\_\_ in the sample solution

 A3 = Peak area of internal standard in the standard solution

 A4 = Peak area of internal standard in the sample solution

 P = Percent purity of Pesticide AI \_\_\_\_\_\_ in the standard

**Result:**

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Sl. No.** | **Name of test** | **Result** | **Unit** | **Method of Analysis**  |
| 1. | Active ingredient |  | % | Method refernce) |
| Remark / Reference : |

|  |  |  |
| --- | --- | --- |
| Analyzed by | Name  |  |
| Dated signature |  |
| Checked by | Name  |  |
| Dated signature |  |